附件1
CT、DR维保服务明细：

	序号
	设备名称
	设备型号
	项目名称
	单位
	数量
	单价（元）
	总价（元）

	1
	GE ACT
	RevolutionACT
	单次人工技术维护保养，调整和校准机器等。
	次
	1
	
	

	2
	飞利浦DR
	DuraDiagnosat F30
	
	次
	1
	
	

	合 计
	
	


维保的技术要求：

	保养检测项目
	执行结果 

	PM Check List (Service Contract Customers Only)
	

	系统 System
	

	检查X射线曝光指示灯Check X-Ray on lights
	

	检查各扫描按钮操作正常Check scan control pushbuttons
	

	检查紧急制动按钮和防触碰开关 (PDU/控制台/扫描架/床)  Verify E-off buttons and Touch Sensor  (PDU/Console/Gantry/Table)
	

	检查系统日志Monitor Sys Logs
	

	校准（如需要）Calibration if needed
	

	图像质量检测IQ check with phantom
	

	备份系统数据 Save System State
	

	检查系统时间的准确性 Check System Clock
	

	操作台 OC
	

	OC整体清洁 General Console Clean (fans/cover air intake screens /filter/vent/console interior / SCSI tower fans and air intake screens)
	

	操作台机柜和台面检查 Visual inspection/check EMC gasket/ check for the loose and worn cables
	

	检查操作台内部风扇Check all fans is operational
	

	检查和清理HP工作站的内部风扇 Check/clean the fans inside host computer (CPU/Video/System/DIMM)
	

	检查和清理DARC/IG的内部风扇和散热器 Check/clean the fans and heat sinks inside DARC/IG (GOC4/LCGOC only)
	

	扫描架 Gantry
	

	检查并清洁扫描窗 Check scan and laser windows
	

	清理通风口/风扇 Clean ventilation holes & check fans
	

	清理滑环粉末 Handling & Removal of Slip Ring Brush Debris
	

	清理/检查滑环和刷子Slip Ring Vacuuming and Brush Inspection
	

	检查数据采集系统/清除灰尘DAS/Duct fan filter cleaning 
	

	检查所有风扇运作情况 Inspect all fans operational
	

	清理DIFB风扇滤网 Clean DIFB Fan Filters (only for Pancake DAS)
	

	检查驱动皮带 Check drive belt for wear
	

	检查前/后倾状态 Tilt speed checking
	

	定位灯检查  Alignment Lights Checks
	

	高压部分 HV System
	

	检查并记录球管扫描次数 Check the number of Tube slices
	

	
	

	检查高压是否存在打火现象 HV Generator No load test
	

	电源分配柜 PDU
	

	盖板、通风口和PDU内部清理Vacuuming cover, vent and PDU interior
	

	检查各器件有无变色/打火痕迹 Visually inspect the PDU electronic components for dis-colorization and arcing damage
	

	紧固各电源线的连接 Tighten all electrical connections for the primary power and sub-system power connections
	

	扫描床 Table
	

	清洁扫描床及扫描床板 Clean the table and cradle
	

	检查扫描床及其附件是否损坏 Check for wear or damage on table cover/head holder/gap sponge/gap between cover and cradle/cradle rail/ body strap/head holder stability/touch sensor
	

	检查液压器件是否漏油 check hydraulic components for oil leak
	

	机械运动部位加润滑油 joint bearing/cylinder/cradle rail grease-up
	


